# EXPERT-CONSENSUS AND EVIDENCE-BASED GUIDELINES FOR THE ASSESSMENT OF FLOW MEDIATED DILATION (FMD) IN HUMANS

Dick HJ Thijssen1,2, Rosa Maria Bruno3, Anke van Mil1,2, Sophie Holder1, Francesco Faita4, Arno Greyling5, Peter Zock5, Stefano Taddei3, John E. Deanfield6, Thomas Luscher7, Daniel J Green1,8, Lorenzo Ghiadoni3

1Research Institute for Sport and Exercise Science, Liverpool John Moores University, Liverpool, United Kingdom

2Radboud Institute for Health Sciences, Department of Physiology,

##### Radboud University Medical Center, Nijmegen, The Netherlands

3Department of Clinical and Experimental Medicine, University of Pisa, Pisa, Italy

4Institute of Clinical Physiology, National Research Council, Pisa, Italy

5Unilever R&D Vlaardingen, Vlaardingen, The Netherlands

6 Institute of Cardiovascular Sciences, University College of London, London, United Kingdom

7 London University Heart Center, Royal Brompton and Harefield Hospitals, London, United Kingdom

8School of Sport Science, Exercise and Health, The University of Western Australia, Crawley, Western Australia, Australia

SHORT TITLE:  Guidelines for flow-mediated dilation

WORD COUNT:  5,925

Word count abstract:  224

Author for correspondence:

# Rosa Maria Bruno

# Department of Clinical and Experimental Medicine, University of Pisa

# c/o Ospedale Santa Chiara, Building #8

# Via Roma 67, 56124 Pisa, Italy

# tel: +39 050 992914 / 050 992711

# fax: +39 050 992409

email: rosamaria.bruno@unipi.it

# abstract

Endothelial dysfunction is involved in the development of atherosclerosis, which precedes asymptomatic structural vascular alterations as well as clinical manifestations of cardiovascular disease (CVD). Endothelial function can be assessed non-invasively using the flow-mediated dilation (FMD) technique. FMD represents an endothelium-dependent, largely nitric oxide (NO)-mediated dilatation of conduit arteries in response to an imposed increase in blood flow and shear stress. FMD is affected by cardiovascular (CV) risk factors, relates to coronary artery endothelial function, and independently predicts CVD outcome. Accordingly, FMD is a tool for examining the pathophysiology of CVD and possibly identifying subjects at increased risk for future CV events. Moreover, it has merit in examining the acute and long-term impact of physiological and pharmacological interventions in humans.

Despite concerns about its reproducibility, the available evidence shows that highly reliable FMD measurements can be achieved when specialised laboratories follow standardised protocols. For this purpose, updated expert-consensus guidelines for the performance of FMD are presented, which are based on critical appraisal of novel technical approaches, development of analysis software, and studies exploring the physiological principles underlying the technique. Uniformity in FMD performance will i. improve comparability between studies, ii. contribute to construction of reference values, and iii. offer an easy accessible and early marker of atherosclerosis that could complement clinical symptoms of structural arterial disease and facilitate early diagnosis and prediction of CVD outcomes.

KEYWORDS: flow-mediated dilation, methodology, cardiovascular disease, brachial artery, vascular function

Cardiovascular disease (CVD) remains the world’s leading cause of morbidity and mortality.1, 2 Development of CVD starts in early childhood and progresses silently for many years,3 ultimately resulting in angina, claudication, myocardial infarction or ischaemic stroke. The endothelium plays a central role in the process of atherosclerosis from the first to the advanced stages. Due to its strategic anatomical position, it is highly responsive to detect various hemodynamic stimuli (e.g. shear stress, circumferential wall strain)4 and regulates vascular tone, growth, adhesion and coagulation in a endocrine-paracrine manner.5 Not surprisingly, impairment of endothelial function precedes development of atherosclerosis, evident by thickening of the arterial wall and plaque formation.6-9 Endothelial dysfunction is also linked to inflammatory changes within atherosclerotic plaques, leading to plaque vulnerability and, subsequently, increased risk for clinical events.3, 10 The intrinsic link between inflammation and endothelial dysfunction is well-established.11, 12

Several (non-) invasive tools to examine arterial endothelial function have been developed. These techniques provide insight into the patho-physiology of atherosclerosis and CVD, but may also predict cardiovascular (CV) events. The present paper discusses flow-mediated dilation (FMD), which represents a popular and widely used non-invasive tool for examining peripheral artery endothelium-dependent dilation, originally introduced in 1992.13 Briefly, this technique adopts ultrasound to examine changes in brachial artery diameter in response to ischaemia (typically 5-minutes), induced by inflating a blood pressure cuff distal from the imaged artery around the forearm to supra-systolic level. Importantly, substantial variation in performing the FMD is present, which impairs its reproducibility and comparison between studies. First, we will briefly discuss techniques utilised to assess endothelial (dys)function, and specifically focus on the clinical value of FMD. Subsequently, we will provide updated expert consensus guidelines for the performance of FMD, which are based on critical appraisal of novel technical advances, development of analysis software, and studies exploring physiological principles underlying the technique.

- Vascular reactivity tests

Endothelial dysfunction, which is characterised by reduced NO availability14, was originally demonstrated by infusing acetylcholine in atherosclerotic coronary arteries.15 Since this technique is invasive and not widely applicable, less invasive studies in the peripheral circulation have been then introduced (supplementary table). A technique with methodological similarities to intra-coronary infusion involves venous plethysmography, typically of the forearm. After arterial cannulation, forearm blood flow changes after infusion of vasoactive substances modulating NO-release are measured. Alternatively, non-invasive techniques include peripheral arterial tonometry (PAT) and FMD, both measuring vasodilator responses after transient ischaemia. Whereas FMD examines macrovascular endothelial function of the brachial artery, PAT may represent a measure of finger microvascular function. Techniques have their specific advantages and disadvantages (supplementary table), as extensively reviewed elsewhere.14 Importantly, non-invasive techniques have been validated against coronary endothelial function, providing the patho-physiological basis for the predictive value of these tests for CVD.16, 17

- Clinical value of FMD

2.1 Association with CVD prognosis

Impaired FMD has been associated with conditions predisposing to atherosclerosis and CVD, representing an early step in developing subclinical target organ damage and late clinical events.18 Brachial FMD is associated with carotid intima-media thickness progression in a population free of CVD7 and in hypertensive, postmenopausal women.19 In a 3-year follow-up study in hypertensive patients, FMD predicted target organ damage progression (including carotid intima-media thickness, pulse wave velocity, albuminuria and left ventricular hypertrophy), even after adjustment for known risk factors.20 Furthermore, impaired FMD represents an independent predictor of in-stent stenosis after single-vessel coronary interventions.21

Several studies have demonstrated the prognostic value of brachial artery FMD for CV events (Table 1).22-25 Meta-analyses indicate a significant 8-13% lower risk of CV events per percent point increase in brachial artery FMD (e.g. from 5% to 6% dilation, Figure 1). This reduction was relevant both in high– and low-risk populations,22, 25 though appeared larger in patients with established CVD.23, 24 Except for one meta-analysis,24 the overall estimated study quality did not significantly influence FMD predictive value.22, 23, 25

One meta-analysis of 14 studies assessed the influence of the site of the cuff position, which was either placed distal (i.e. forearm) or proximal (i.e. arm) from the imaged artery (Table 1).26 This study found that the predictive capacity of the FMD using distal cuff position did not differ from using FMD with proximal cuff positioning. This was confirmed in a later meta-analysis that included a larger number of studies.24

2.2 Incremental value of FMD over traditional cardiovascular risk factors.

FMD has failed to demonstrate an added value to classical CV risk factors in terms of discrimination and net reclassification.27 In the Multi-Ethnic Study of Atherosclerosis (MESA) cohort of intermediate risk individuals, FMD did not increase discriminative power or net reclassification index of risk compared to the Framingham Risk Score.28 However, the low reproducibility of the FMD in the MESA study, in which the intra-class correlation coefficient was 0.5428 (worse than other studies in healthy volunteers 29, 30 and CVD patients30, 31), may have contributed to these negative findings. Alternatively, FMD may simply reflect compound risk burden that impacts upon vessel function and hence may not provide incremental risk prediction. A recent position paper by the European Society of Cardiology Working Group on Peripheral Circulation32 concluded that FMD is principally a valuable research tool. Moreover, authors stated that, partly because poor standardisation between laboratories and lack of guideline adherence, FMD is currently not recommended for the assessment of CV risk, with evidence grading remaining at III, in accordance with other guidelines.33, 34

2.3 Clinical value of long-term changes in FMD.

Longer-term improvement in FMD may have a prognostic implication.35 Therefore, assessment of FMD in interventional trials could represent a surrogate endpoint, especially since FMD responds rapidly to therapies, allowing identification and selection of new drugs or bioactive substances.36 The lack of improvement in FMD might identify “non-responder” patients, who might be suitable for more intensive or new therapeutic approaches. For example, improvement in FMD following 6-months of antihypertensive therapy was associated with a more favourable prognosis for CV events in hypertensive postmenopausal women.37 Furthermore, persistent impairment in FMD after optimised risk reduction therapy represents an independent predictor of CV events in patients with coronary artery disease.38 Finally, a persistently impaired FMD, after optimised therapy in heart failure, independently predicted cardiac events.39 These studies support the potential use of repeated assessment of FMD, rather than a single measure, to predict future CV events.

2.4 Potential value of short-term changes in FMD.

Many studies have examined short-term changes in FMD, ranging from a few hours, to several days, and up to months. In these randomised trials, FMD was selected as the primary outcome measure to investigate the potential protective impact of (non-)pharmacological therapies.36, 40-42 Increase in FMD in response to (non-)pharmacological substances may translate to long-term improvements in endothelial function and protection against atherosclerosis and/or CVD. The relative simplicity of the design and the ability to perform reproducible, repeated measurements explain the popularity of examining changes in FMD in these studies.14

In summary, a biomarker is only useful for CV risk prediction when it adds incremental information to traditional CV risk factors, being independently associated with outcomes, and also improving discrimination, calibration and net reclassification of risk, especially in those at intermediate risk.43 Current evidence on these issues pertaining to FMD is incomplete. In part, this seems related to the diversity between studies in adherence to expert-consensus guidelines. Moreover, reference values are yet to be established and are necessary for widespread applicability. Nonetheless, FMD has proven beneficial when examining (short-term) changes to evaluate the effects of (non-)pharmacological interventions.

3. Ultrasound assessment: technical requirements and common mistakes

3.1 Diameter and velocity assessment.

FMD is typically examined in the brachial artery (diameter 3-5 mm). Longitudinal images of the brachial artery are taken using high-resolution B-mode ultrasound, usually with an ultrasound probe of 7.5-12 MHz. Studies have also examined smaller (e.g. radial artery: diameter 1.5-3 mm) or larger sized arteries (e.g. superficial femoral artery: diameter 5-7 mm, popliteal artery: diameter 4-6 mm). Changes in resolution may be required to optimize B-mode images in arteries that lie at differing depths. A key challenge with B-mode imaging is to identify clear vascular boundaries, i.e. the double lines of Pignoli (Figure 2).44, 45,46 A reasonable approach involves obtaining an image characterised by a single echogenic layer, representing the brachial artery wall-lumen interface. Since diameter is a key determinant of FMD, the modality of measurement applied may influence results and, under all circumstances, should be kept the same throughout the experiment. Special attention should be paid to perform adequate scanning of the baseline diameter. More specifically, tangential scanning is a common error and results in underestimation of the true brachial artery diameter. Recent technological advances, which adopt an H-shaped probe capturing 2 short-axis and 1 long-axis for automatic probe position correction, may overcome this crucial limitation.47

Simultaneous live acquisition of the pulsed-wave Doppler velocity is recommended, given the importance of shear stress as the eliciting stimulus for dilation.48-50 An important limitation of this approach is that the same transducer is employed to detect Doppler velocity and arterial diameter, which have competing requirements for optimal data acquisition.  B-mode echoes are of greater intensity with perpendicular insonation of the ultrasound beam (90°), whereas optimal pulse-waved Doppler signals require parallel incidence with the direction of blood flow (ideally 0°, Figure 2).51-53 Therefore, a compromise must be reached to uphold fundamental principles and assumptions for both modalities.53 The error linked to incorrect estimation of insonation angle increases exponentially with angles greater than 60º.51-53 Therefore, we recommend an insonation angle of <60-70º (which is the best achievable angle with standard ultrasound machines), which should be kept constant and reported in the manuscript.

Finally, the sample volume used for pulsed-wave Doppler should be considered. Since measurements settings of the width of the sample volume vary between laboratories, and impact upon the analysis of the velocity signal (see section 3.2), we recommend to maintain consistency in methods within studies, and in particular for repeated tests within subjects.

3.2 Analysis of the velocity signal.

Duplex ultrasound-derived velocity and diameter data are both required for calculating shear stress, the eliciting stimulus for artery dilation during FMD. Although the increase in shear stress is the physiological stimulus for dilation, most studies present shear rate, as it is assumed that blood viscosity does not differ between participants and/or groups.50, 54, 55 Since no uniformity exists in the calculation of shear rate,56 calculations must be described in the manuscript. Two different formulas have been suggested, according to the sample volume size and placement:

a) Shear rate = 8 × mean blood velocity / internal diameter; for large, centred sample volume;

b) Shear rate = 4 × mean blood velocity / internal diameter; for small, centred sample volume.

A clear description must also be provided for the calculation of velocity. Mean velocity can be estimated by taking half of the peak velocity (i.e. fastest moving blood cells in the centre of the vessel) or the intensity weighted mean velocity (i.e. mean velocity from all Doppler shifts across the vessel).52 Although both methods seem valid, they cannot be used interchangeably.57

3.3 Maintaining the optimal image: probe holder versus hand-held.

Another common mistake is inconsistency in scanning the same portion of the artery during repeated measurements, or throughout a single FMD. Therefore, anatomical landmarks should be identified and recorded and/or the distance between the elbow crease and ultrasound probe should be recorded (or photographed). The position of the subject (related to the arm and hand), but also cuff size and position (see 4.2.1) should be also recorded and standardised.58

To further ensure optimal image throughout the FMD-procedure, a probe-holding device can be useful.29, 58, 59 A stereotactic adjustable probe-holding device allows adjustment of probe position during the test, ensuring that the same scan is maintained throughout the study.18 Subject movement can also be compensated for when the probe hand-held, assuming an experienced sonographer. In a recent systematic review of studies that examined the reproducibility of the FMD, the use of a probe-holding device was associated with a higher FMD reproducibility.60 However, another recent study pooling repeated measurements from different centres found that reproducibility is not affected by a probe-holder device, when experienced laboratories perform FMD following strict guidelines.30 It cannot be excluded that other differences between laboratories may explain these findings. Nevertheless, a probe-holding device is a cheap, easily obtainable tool. Because it may facilitate FMD testing for inexperienced operators and may increase accuracy, a probe-holder device is recommended for laboratories with less experience in FMD measurements.

4. Practical considerations, study protocol and data analysis

4.1 Subject preparation and environmental influences.

Subject preparation is important in the valid and reproducible assessment of the FMD, with many studies failing to control for these factors and/or failing to report on their procedures in sufficient detail. Several subject-related factors can influence FMD, including food, alcohol, smoking, supplements, drugs, physical activity, and mental stress.61-64 While some factors directly impact stimulated NO-release, others, such as acute physical exercise,62 modify baseline vasomotor tone (and therefore baseline diameter).4 To minimize the effect of these confounding factors, it is strongly recommended that prior to FMD examination, subjects are fasted (>6 h), avoid exercise (>24 h), and refrain from caffeine, vitamin C, polyphenols, alcohol and supplements known to affect the cardiovascular system for a consistent period of time (typically >12 hours) (Table 2). Smokers must refrain from smoking for a standardised period (preferably >6h). When examining patient groups taking drugs, we recommend waiting 4 times the half-life of the drug. If drug intake cannot be avoided, examination should be performed after a consistent and standardized time.59 Since mental stress affects FMD,64, 65 testing must be performed in a quiet, temperature-controlled room after a standardised period of supine rest of >10-15 minutes. Acute intense mental stress can cause prolonged (up to 90 min) impairment in FMD66, but its confounding effect can be hardly controlled or eliminated. Premenopausal women should be examined in a standardised phase of the menstrual cycle, since hormonal changes can affect FMD.67 Vascular function in humans, including that assessed by FMD,68 demonstrates a diurnal variation. This implies that the time of the day of FMD assessment should be standardised and reported, especially when performing repeated measures within subjects. Other environmental factors that may affect FMD are outdoor temperature,69 seasonality70 and air pollution.71 However, because of practical reasons, these factors should not be considered as limiting when the ideal conditions cannot be achieved.

4.2. Study protocol

4.2.1 Protocol: cuff position.

Placement of the occlusion cuff importantly alters the magnitude,72 duration,72, 73 nature74 and possibly the clinical relevance26 of the dilator response. Originally, Celermajer et al.13 examined brachial artery FMD with the occlusion cuff placed around the forearm, distal to the ultrasound probe (Figure 2). Standardising cuff position is of crucial importance for valid comparisons between studies, since occlusion of a smaller volume (i.e. placement towards the wrist) or larger volume of tissue (i.e. cuff occlusion around the upper arm) leads to either a markedly attenuated diameter58 or an enhanced72, 75 shear stress response.

Some have supported the placement of the occlusion cuff above the imaged artery,73 with the main argument that the larger and sustained dilator response76 achieved by occluding a larger area may improve discriminative power to detect subtle differences between populations or within subjects. One analysis found no significant differences in prognostic value between studies that adopted distal versus proximal cuff position.26 Furthermore, some limitations of proximal cuff inflation must be considered. First, scanning the brachial artery during this procedure is challenging, since the artery can collapse and/or tissue movements can distort image quality. Secondly, the dilator response with proximal cuff position is largely mediated through vasoactive substances other than NO.61 A meta-analysis74 found that the dilation in response to distal cuff occlusion is ~70% NO-mediated, whereas proximal cuff placement is only ~30% NO-mediated. These differences between studies are confirmed by direct comparisons of distal versus proximal cuff position in the same individual.73 This is mechanistically important, because several studies specifically aim to study the NO-dependent pathway, given its established importance in the process of atherosclerosis. Thus, strict standardization of distal cuff occlusion (i.e. below the imaged artery) is recommended to ensure maximal dependence of the dilator response on endothelium-derived NO.

4.2.2 Protocol: baseline diameter.

The FMD-response is typically expressed as the relative change in post-hyperaemia diameter from baseline. As a direct consequence and limitation, the FMD is a function of the degree of responsiveness to stimuli, but also of baseline vasomotor tone and structural remodeling. Correct assessment of baseline diameter is therefore of crucial importance and implies the removal of potential factors influencing both baseline diameter and vasomotor tone (see section 4.1), but also practical (3.1+this section) and statistical procedures (4.4.2+this section) linked to baseline diameter.

Most studies use the pre-inflation diameter to calculate the FMD, in line with the classic approach13 and previous FMD-guidelines,58, 59, 61 It is recommended that pre-inflation diameter should be recorded for 1-minute, with a minimum of 30-seconds.61 The end-of ischemia diameter has been also proposed as the baseline diameter for FMD calculation. However, some studies found that the “low-flow” state during distal cuff occlusion leads to artery constriction, though this phenomenon appears to be evident mostly in the radial artery,77-79 rather than in the brachial artery.49, 80-82 Conversely, one study showed that brachial artery diameter during cuff inflation was significantly larger than that assessed pre-inflation, leading to a lower FMD, an effect not present in older groups.83 Based on these between-group differences, we recommend using pre-occlusion diameters as the baseline value so that the differential impacts of cuff occlusion on the baseline artery diameter are avoided.

Because FMD is defined as increase in diameter after occlusion, baseline diameter is strongly and inversely related with FMD.13, 84, 85 This relation has several important implications for comparing groups that differ in baseline diameter (e.g. age, CVD),86, 87 and for interpretation of studies with interventions that may affect resting diameter (e.g. exercise training).4 Atkinson et al. found that the FMD% does not accurately scale for inter-individual differences in baseline diameter, but overestimated endothelial function for low baseline diameters and vice versa.88, thus proposed to adopt allometric scaling, a statistical method to account for the relationship between baseline and peak diameter.84 These approaches may be relevant, particularly when large differences in baseline diameter are present between or within subjects. The use of allometric scaling, however, is limited by its inability to adjust FMD values at an individual level.

4.2.3 Protocol: duration and magnitude of cuff occlusion.

The original 5-minute duration of cuff occlusion remains the most frequently used protocol. Shorter periods lead to negligible dilation,89 whilst longer periods lead to larger blood flow and diameter responses, but may be less tolerable for volunteers.72 Since the nature of dilation changes with longer periods of occlusion,75 we recommend using a 5-minute occlusion period.

Whilst a majority of studies use a pre-defined cuff occlusion pressure (typically between 200 and 300 mmHg), it is important that cuff pressure should exceed >50 mmHg above systolic pressure to prevent arterial inflow. In our experience, the procedure is generally well tolerated and drop-out due to discomfort is unusual, though no studies are available in literature specifically dealing with this aspect. A common-sense recommendation to minimise subject discomfort is to use a smaller-sized (5-cm), paediatric cuff instead of a standard (12-cm) cuff. In addition, providing clear information about the study protocol and FMD-procedure is essential.

4.2.4 Protocol: post-deflation measurements.

In the first studies using the FMD approach, peak diameter was examined using static frames at or around 60-seconds post-deflation.13 During the past decade, several studies found that this approach (or using pre-determined time-windows) can significantly underestimate the true peak dilation.90 More importantly, the timing of the peak dilation may differ between groups90 or after interventions.90-93 Therefore, to ensure successful capture of the true peak brachial artery diameter, guidelines support continuous examination up to 180-seconds post-deflation.

Endothelial function is increasingly examined in conduit arteries other than the brachial artery. However, the nature of the diameter responses in other arteries may differ from that in the brachial artery. For example, the content of endothelial NO synthase is heterogeneous throughout the arterial tree.94 This may impact upon the relative contribution of NO to the dilator response. In addition, lower limb arteries show a significantly delayed peak dilation compared to those in the arms.95 This means that the 3-minute post-deflation time window may not be long enough for all arteries. These potential differences must be considered when performing and interpreting the FMD in other conduit arteries.

4.3 Identification of the peak diameter.

Early studies applied ECG-gated analysis to identify peak diameter.59 This approach assesses the diameter at the onset of R-waves (i.e. end-diastole), which limits the influence of arterial compliance on the assessment of the diameter. In the past decade, driven by the technical possibility and recognition of its relevance, studies typically adopted continuous data analysis of the diameter across the cardiac cycle. Continuous data analysis shows good agreement with ECG-gated analysis96, 97 and is more time-efficient.97 The accuracy and availability of this procedure makes ECG-gating no longer mandatory for determining peak artery diameter.

4.4. Shear stimulus

4.4.1 Shear stimulus: importance for FMD

Although the terminology of FMD suggests that dilation is flow-mediated, physiological work demonstrates that it is rather mediated through the post-deflation increase in shear stress.48 Studies that adopted within-subject manipulation of the post-deflation shear stress stimulus revealed that the relation between shear stress and dilation is dose-dependent.48, 49, 98, 99 This fits with landmark studies100, 101 demonstrating the importance of shear stress in mediating endothelium-dependent dilation. Pyke and Tschakovsky49 suggested that the total, rather than peak shear stress, is more important in mediating conduit artery dilation.49 This finding was later confirmed by others.50 These studies show physiological and mechanistic basis for the continuous recording of both conduit artery diameter and velocity (Figure 1).

4.4.2 Shear stimulus: normalization for FMD stimulus

Based on the importance of shear stress as the eliciting stimulus, and also the assumption that within-/between-subject variation in FMD-responses relate to the magnitude of reactive hyperaemia, studies have used different approaches to account for the shear stress stimulus. An early approach involved ‘simple’ ratio normalization by dividing the FMD by the shear rate stimulus.50, 81, 102 However, this approach tends to violate important statistical assumptions being (i) the relationship between both parameters is linear, (ii) the intercept for the regression slope of this relationship is zero, (iii) data (including residuals) are normally distributed, (iv) variances are similar between groups and (v) the ratio does not lead to spurious correlations with other variables.103 Whilst acceptable-to-good relation between the shear rate stimulus and FMD is present within subjects,48-50, 89, 99, 104 studies examining this relation between subjects found a weak95 or absent105 relation. Whilst this does not invalidate the role of shear stress as the dilator stimulus, it indicates that FMD variability cannot be simply controlled for by ratio normalization.103

Studies have also explored other strategies to statistically correct the FMD-response for the shear stress stimulus, such as including shear rate as a covariate in an analysis of covariance (ANCOVA).106-108 ANCOVA results, however, could be misleading if the covariate (i.e. shear rate), is related to the independent and/or outcome variable (e.g. age, sex, baseline diameter and FMD). However, since calculation of shear rate importantly depends on baseline diameter, shear rate normalization may not be needed if the FMD is scaled properly to baseline diameter using the previously mentioned allometric scaling.109

Taken together, it is currently not clear which statistical strategy is preferred to account for the eliciting shear rate stimulus in the FMD response. Ratio-normalization seems statistically flawed, whilst too many questions remain present around validity, practicality and potential clinical importance of statistical correction for the shear rate stimulus. It is recommended to report the relevant shear rate stimulus, i.e. shear area-under-the-curve up to the peak diameter. Reporting the shear stress stimulus may also be clinically relevant because the post-occlusion hyperaemia is strongly related to CV risk, and may have prognostic value for future CV events.110-113

4.5 Analysis: Blinding, edge-detection and wall-tracking

Early FMD research relied on a manual approach to assess diameter change through visual inspection and placement of calipers.13, 114 This method is highly operator-dependent, carries risk of observer error, and is time-consuming.115-119 Recent years, software systems for automatic diameter measurement have been developed. This represents a time-efficient approach, and has limited operator-related errors and bias. More importantly, intra-observer variation is significantly lower with automated analysis of multiple measurements compared to the classic manual technique.46, 115, 116, 118-120

Automatic wall tracking algorithms belong to two main categories according to the input signal: radio-frequency (RF) backscattering or grey scale speckle pixel intensities. In general, RF-based algorithms are mono-dimensional with raw data processed in a single vertical line of view, while speckle tracking is applied on large bi-dimensional regions of interest. Previously, the RF-system was considered more accurate due to the higher axial spatial resolution, but this comes at the cost of assessing a single arterial diameter site (analogous to the use of a single caliper dimension). Pixel density systems allow for multiple diameter assessments within a given frame, whilst recent speckle tracking systems have also overcome the spatial resolution limitation.121

Adoption of automatic software overcomes methodological issues linked to the use of manual calipers. Since maximal brachial vasodilation is 0.1-0.4 mm and the typical resolving power of manual calipers of 0.1 mm, the manual approach is relatively limited and likely to introduce significant within- and between-observer error. Manual measurements also have a low number of samples per time frame, whilst a more detailed time course of changes in diameter is desirable, especially to identify the true peak diameter.90 Other potential advantages offered by automatic systems is that real-time analysis allows instantaneous feedback on the quality of the scan and allows for rapid adjustments in response to movements of the patient/probe. Presence of an optimal distension signal, which is characterised by a steep systolic rise followed by a dicrotic notch, suggests correct placement of the probe. This real-time feedback may help the sonographer to optimise image quality. Another potential advantage is to immediately identify technical failure of the measurement, with the possibility to repeat the FMD after an adequate resting time.

FMD studies with manual versus automated analysis are hardly comparable, due to differences in estimations of diameter and shear rate.116 Therefore, validated, accurate and reproducible edge-detection and wall tracking systems must be used. An additional advantage of automated systems is their ability to use both ECG-gated and non-gated ultrasound images without affecting the FMD results.96 This can help to keep the experimental setup as simple as possible, for example by adopting a cheaper entry-level ultrasound device without ECG synchronization capability. Software systems for automatic FMD evaluation should be usable and user-friendly, but also should provide options to view and record different parameters (i.e. shear rate). Furthermore, the high throughput of data, ease of use and other technical features (e.g. real-time feedback, overlay of edge detector output with B-Mode) may help adequate training of the operator. Finally, any new automatic system must be validated in appropriate populations before introducing its use in laboratories.45, 122

4.6.Summary and future directions.

In conclusion, strong evidence supports the following methodologies: placement of a lower cuff occlusion and use of continuous diameter monitoring by automatic edge-detections systems. Furthermore, the general study protocol, subject preparation and lab requirements are well defined. It remains uncertain whether and how to correct FMD for the shear stimulus and/or baseline diameter. The patho-physiological significance and clinical relevance of additional parameters, such as flow-mediated constriction and reactive hyperaemia as a marker of microcirculatory function, are still matter of debate and require further investigation.

5. Reproducibility of FMD measurements

Several studies have examined FMD reproducibility. However, these have been performed with widely varying adherence to FMD-guidelines. After pooling the results of 7 centers that strictly followed FMD-guidelines, including optimal training of operators and monitoring, Ghiadoni et al. found that FMD in healthy volunteers is highly reproducible among centers (coefficient of variation 11.6-16.1%).29 Data from the dal-VESSEL trial also showed reproducible results in 19 laboratories after standardised training in patients with CVD (coefficient of variation 15.8-17.5%).31 More recently, Greyling and colleagues designed a tool to evaluate adherence to FMD expert consensus guidelines.123 Subsequently, a meta-analysis was performed to assess the relation between degree of adherence and reproducibility of the FMD.60 Not surprisingly, stricter adherence to guidelines was related with markedly less measurement error in brachial artery FMD. This work highlights that strictly following contemporary guidelines leads to highly reproducible results.

5.1 Factors influencing reproducibility: group and methodology differences

Guidelines for FMD measurement cannot cover and eliminate all sources of measurement variations and error. Understanding of these factors is important for designing studies on FMD (e.g. sample size calculations). Van Mil et al. recently performed analysis of 672 repeated FMD measurements that were collected after strictly following contemporary guidelines.30 High reproducibility was found in healthy subjects (9.3%), whilst regression models several factors that contributed to variations in reproducibility (see below).

Subject characteristics.  Subjects with older age, hypertension, or dyslipidaemia show a larger variability in FMD.29-31 This supports earlier findings in that CV risk factors impair FMD reproducibility.124 An explanation for the larger variation in high-risk subjects may relate to the lower baseline FMD% typically observed in these populations.58, 125, 126 Indeed, Van Mil et al. demonstrated that a lower FMD% explained part of the higher variability in FMD with older age and dyslipidaemia.124 Differences in vascular structure and compliance may also contribute to this higher variability127 and finally, not all subjects may be equally susceptible to the effects of CV risk factors at the endothelial level.

Methodological characteristics (time between measurements). Longer time periods between subsequent FMD measurements leads to larger variation.29, 30 Nonetheless, when strictly following guidelines and controlling for potential physiological factors,29 it is possible to detect FMD changes, induced by interventions, over time with a relatively low sample size. Current data on within- and between-subject variability justify the use of FMD as an outcome measure in short-to-medium-term studies to evaluate (non)pharmacological interventions in humans.

Methodological factors (lab experience). Possibly the most important methodological source of variability in the FMD is proper training of the operators. Guidance and training by expert colleagues seems crucial since lower variability is present in more experienced laboratories. 30 Although the need of experience in scanning is mentioned by previous guidelines,58, 59, 61 it is difficult to quantify “proper training”. No study specifically aimed at quantifying the duration of the training necessary to achieve reproducible FMD measurements, like what was done for carotid intima-media thickness.128 Corretti et al. suggested that >100 scans/year are required to maintain competency.59 However, there was no mention of the training required before performing FMD measurements. Two independent consortia adopted a formal training for sonographers: sonographers were qualified for measurements when the CV for repeated scans were <2% for brachial artery diameter and <15% for FMD.29, 31 Since FMD variability seems larger in individuals with CV risk,29, 31 operator training might be longer depending on the study population enrolled. Both aforementioned studies showed excellent reproducibility among centres, including heterogeneous populations, suggesting that specialised training is achievable and required to achieve a minimum level of variability. We therefore recommend that all studies on FMD should report details on preparation and training of the sonographers, along with the observed variability in their FMD measurements.

6. Endothelium-independent dilation

6.1 Pathophysiology and clinical relevance

The degree of FMD can also be influenced by the functionality of vascular smooth muscle (VSM) cells.129 Any functional defect in VSM cells reduces their capacity to respond to NO, whilst a compensatory increased VSM response may be present in endothelial dysfunction.130, 131 In any case, it is recommended that the studies on endothelium-dependent vasodilation also determine the extent of any coexistent endothelium-independent dysfunction. Although this recommendation was already present in previous FMD-guidelines,59 a standardised protocol has not been clearly defined yet.

Clinical determinants of VSM function are not fully understood. In a large Japanese cohort, brachial artery dilation to sublingual glyceryl trinitrate (GTN) was correlated with most classical CV risk factors and represented an independent determinant of FMD.132 Conversely, a meta-analysis reported micro-, but not macrovascular, impairment of VSM in type 2 diabetic patients.133 Since the brachial artery was studied in only 12 out of the 31 studies, this subgroup analysis is likely underpowered. Furthermore, some studies found a relation between brachial artery GTN-responses and lack of nocturnal blood pressure fall134, coronary artery calcium,135 and microalbuminuria.136 Although GTN-responses show a relationship with Framingham Risk score,132 only few prospective studies investigated the relation between this response and CVD events.137, 138 Thus, the prognostic value of the GTN-response remains mostly inconclusive.137, 138 Finally, the doses typically used in these studies induce near-maximal dilation,72 hence the GTN-responses may be a structural index rather than a functional one.

6.2 Dose, technique, and administration of exogenous nitrates

Many research centers use sublingual GTN 400 mcg, the lowest marketed dose in many countries. Administration of GTN via a sublingual spray leads to a larger and faster response than by tablet administration.139 Lower dosages (25 mcg) of GTN were later introduced140 and recommended in the 2005 ESH-statement.141 The rationale for lower dosages is the induction of less extreme dilatation, a faster return to baseline diameter (~20-min), and lower risk of side effects. In healthy volunteers, 8-35 mcg GTN is estimated to induce a 4%-10% dilation, without changes in blood pressure and heart rate.132, 142 Furthermore, 25-mcg GTN did not cause any significant increase in sympathetic signalling directed to muscle vasculature, evaluated by microneurography.143 It has been suggested that higher GTN doses might be useful to assess maximum vasodilating capacity, thus providing information about structural alterations. However, maximal vasodilation cannot conceivably be obtained in these conditions because of counter-regulatory sympathetic vasoconstriction induced by the high doses of GTN. Although a low-dose GTN is recommended, this approach may not always be feasible, since low-dose GTN for oral use is not commercially available for most countries.

GTN should be administered >10-min after FMD testing to ensure return to baseline diameter. Furthermore, continuous diameter monitoring is recommended, because the timing of peak dilatation varies between subjects.144 The peak vasodilation response to GTN usually occurs between 4-/5-min, both with high144, 145 and low GTN doses.140 Therefore, monitoring should be performed >5-min after GTN administration.

7. Summary  Conclusions

FMD provides valuable and independent prognostic information. Unfortunately, different methodological approaches importantly limit its validity, comparability and its potential use as a clinical and physiological research tool. Indeed, adherence to guidelines and appropriate operator training improve FMD variability. Therefore, performing and reporting FMD according to state-of-the-art guidelines (see Table 2) is crucial to ensure valid conclusions and clinical evaluation. The application of state-of-the art methodology allows examining whether the FMD correctly reclassifies individuals, adding incremental information to traditional CV risk factors.  Finally, strict adherence to guidelines will also contribute to reference values, which will further improve the clinical applicability of FMD.

REFERENCES:

1.	Townsend N, Nichols M, Scarborough P, Rayner M. Cardiovascular disease in Europe--epidemiological update 2015. Eur Heart J 2015;36(40):2696-705.

2.	Timmis A, Townsend N, Gale C, Grobbee R, Maniadakis N, Flather M, Wilkins E, Wright L, Vos R, Bax J, Blum M, Pinto F, Vardas P, Group ESCSD. European Society of Cardiology: Cardiovascular Disease Statistics 2017. Eur Heart J 2018;39(7):508-579.

3.	Ross R. The pathogenesis of atherosclerosis: a perspective for the 1990s. Nature 1993;362(6423):801-9.

4.	Green DJ, Hopman MT, Padilla J, Laughlin MH, Thijssen DH. Vascular Adaptation to Exercise in Humans: Role of Hemodynamic Stimuli. Physiol Rev 2017;97(2):495-528.

5.	Vanhoutte PM, Shimokawa H, Tang EH, Feletou M. Endothelial dysfunction and vascular disease. Acta Physiol (Oxf) 2009;196(2):193-222.

6.	Juonala M, Viikari JS, Laitinen T, Marniemi J, Helenius H, Ronnemaa T, Raitakari OT. Interrelations between brachial endothelial function and carotid intima-media thickness in young adults: the cardiovascular risk in young Finns study. Circulation 2004;110(18):2918-23.

7.	Halcox JP, Donald AE, Ellins E, Witte DR, Shipley MJ, Brunner EJ, Marmot MG, Deanfield JE. Endothelial function predicts progression of carotid intima-media thickness. Circulation 2009;119(7):1005-12.

8.	Glowinska-Olszewska B, Tolwinska J, Urban M. Relationship between endothelial dysfunction, carotid artery intima media thickness and circulating markers of vascular inflammation in obese hypertensive children and adolescents. J Pediatr Endocrinol Metab 2007;20(10):1125-36.

9.	Kobayashi K, Akishita M, Yu W, Hashimoto M, Ohni M, Toba K. Interrelationship between non-invasive measurements of atherosclerosis: flow-mediated dilation of brachial artery, carotid intima-media thickness and pulse wave velocity. Atherosclerosis 2004;173(1):13-8.

10.	Ross R. Atherosclerosis--an inflammatory disease. N Engl J Med 1999;340(2):115-26.

11.	Hurlimann D, Forster A, Noll G, Enseleit F, Chenevard R, Distler O, Bechir M, Spieker LE, Neidhart M, Michel BA, Gay RE, Luscher TF, Gay S, Ruschitzka F. Anti-tumor necrosis factor-alpha treatment improves endothelial function in patients with rheumatoid arthritis. Circulation 2002;106(17):2184-7.

12.	Bilsborough W, Keen H, Taylor A, O'Driscoll GJ, Arnolda L, Green DJ. Anti-tumour necrosis factor-alpha therapy over conventional therapy improves endothelial function in adults with rheumatoid arthritis. Rheumatol Int 2006;26(12):1125-31.

13.	Celermajer DS, Sorensen KE, Gooch VM, Spiegelhalter DJ, Miller OI, Sullivan ID, Lloyd JK, Deanfield JE. Non-invasive detection of endothelial dysfunction in children and adults at risk of atherosclerosis. Lancet 1992;340(8828):1111-5.

14.	Flammer AJ, Anderson T, Celermajer DS, Creager MA, Deanfield J, Ganz P, Hamburg NM, Luscher TF, Shechter M, Taddei S, Vita JA, Lerman A. The assessment of endothelial function: from research into clinical practice. Circulation 2012;126(6):753-67.

15.	Ludmer PL, Selwyn AP, Shook TL, Wayne RR, Mudge GH, Alexander RW, Ganz P. Paradoxical vasoconstriction induced by acetylcholine in atherosclerotic coronary arteries. N Engl J Med 1986;315(17):1046-51.

16.	Takase B, Hamabe A, Satomura K, Akima T, Uehata A, Ohsuzu F, Ishihara M, Kurita A. Close relationship between the vasodilator response to acetylcholine in the brachial and coronary artery in suspected coronary artery disease. Int J Cardiol 2005;105(1):58-66.

17.	Takase B, Uehata A, Akima T, Nagai T, Nishioka T, Hamabe A, Satomura K, Ohsuzu F, Kurita A. Endothelium-dependent flow-mediated vasodilation in coronary and brachial arteries in suspected coronary artery disease. Am J Cardiol 1998;82(12):1535-9, A7-8.

18.	Charakida M, Masi S, Luscher TF, Kastelein JJ, Deanfield JE. Assessment of atherosclerosis: the role of flow-mediated dilatation. Eur Heart J 2010;31(23):2854-61.

19.	Rossi R, Nuzzo A, Olaru AI, Origliani G, Modena MG. Endothelial function affects early carotid atherosclerosis progression in hypertensive postmenopausal women. J Hypertens 2011;29(6):1136-44.

20.	Yang Y, Xu JZ, Wang Y, Tang XF, Gao PJ. Brachial flow-mediated dilation predicts subclinical target organ damage progression in essential hypertensive patients: a 3-year follow-up study. J Hypertens 2014;32(12):2393-400; discussion 2400.

21.	Patti G, Pasceri V, Melfi R, Goffredo C, Chello M, D'Ambrosio A, Montesanti R, Di Sciascio G. Impaired flow-mediated dilation and risk of restenosis in patients undergoing coronary stent implantation. Circulation 2005;111(1):70-5.

22.	Inaba Y, Chen JA, Bergmann SR. Prediction of future cardiovascular outcomes by flow-mediated vasodilatation of brachial artery: a meta-analysis. Int J Cardiovasc Imaging 2010;26(6):631-40.

23.	Ras RT, Streppel MT, Draijer R, Zock PL. Flow-mediated dilation and cardiovascular risk prediction: A systematic review with meta-analysis. Int J Cardiol 2013;168(1):344-51.

24.	Matsuzawa Y, Kwon TG, Lennon RJ, Lerman LO, Lerman A. Prognostic Value of Flow-Mediated Vasodilation in Brachial Artery and Fingertip Artery for Cardiovascular Events: A Systematic Review and Meta-Analysis. J Am Heart Assoc 2015;4(11).

25.	Xu Y, Arora RC, Hiebert BM, Lerner B, Szwajcer A, McDonald K, Rigatto C, Komenda P, Sood MM, Tangri N. Non-invasive endothelial function testing and the risk of adverse outcomes: a systematic review and meta-analysis. Eur Heart J Cardiovasc Imaging 2014;15(7):736-46.

26.	Green DJ, Jones H, Thijssen D, Cable NT, Atkinson G. Flow-mediated dilation and cardiovascular event prediction: does nitric oxide matter? Hypertension 2011;57(3):363-9.

27.	Peters SA, den Ruijter HM, Bots ML, Moons KG. Improvements in risk stratification for the occurrence of cardiovascular disease by imaging subclinical atherosclerosis: a systematic review. Heart 2012;98(3):177-84.

28.	Yeboah J, McClelland RL, Polonsky TS, Burke GL, Sibley CT, O'Leary D, Carr JJ, Goff DC, Greenland P, Herrington DM. Comparison of novel risk markers for improvement in cardiovascular risk assessment in intermediate-risk individuals. JAMA 2012;308(8):788-95.

29.	Ghiadoni L, Faita F, Salvetti M, Cordiano C, Biggi A, Puato M, Di Monaco A, De Siati L, Volpe M, Ambrosio G, Gemignani V, Muiesan ML, Taddei S, Lanza GA, Cosentino F. Assessment of flow-mediated dilation reproducibility: a nationwide multicenter study. J Hypertens 2012;30(7):1399-405.

30.	van Mil AC, Greyling A, Zock PL, Geleijnse JM, Hopman MT, Mensink RP, Reesink KD, Green DJ, Ghiadoni L, Thijssen DH. Impact of volunteer-related and methodology-related factors on the reproducibility of brachial artery flow-mediated vasodilation: analysis of 672 individual repeated measurements. J Hypertens 2016;34(9):1738-45.

31.	Charakida M, de Groot E, Loukogeorgakis SP, Khan T, Luscher T, Kastelein JJ, Gasser T, Deanfield JE. Variability and reproducibility of flow-mediated dilatation in a multicentre clinical trial. Eur Heart J 2013;34(45):3501-7.

32.	Vlachopoulos C, Xaplanteris P, Aboyans V, Brodmann M, Cifkova R, Cosentino F, De Carlo M, Gallino A, Landmesser U, Laurent S, Lekakis J, Mikhailidis DP, Naka KK, Protogerou AD, Rizzoni D, Schmidt-Trucksass A, Van Bortel L, Weber T, Yamashina A, Zimlichman R, Boutouyrie P, Cockcroft J, O'Rourke M, Park JB, Schillaci G, Sillesen H, Townsend RR. The role of vascular biomarkers for primary and secondary prevention. A position paper from the European Society of Cardiology Working Group on peripheral circulation: Endorsed by the Association for Research into Arterial Structure and Physiology (ARTERY) Society. Atherosclerosis 2015;241(2):507-32.

33.	Goff DC, Jr., Lloyd-Jones DM, Bennett G, Coady S, D'Agostino RB, Gibbons R, Greenland P, Lackland DT, Levy D, O'Donnell CJ, Robinson JG, Schwartz JS, Shero ST, Smith SC, Jr., Sorlie P, Stone NJ, Wilson PW, Jordan HS, Nevo L, Wnek J, Anderson JL, Halperin JL, Albert NM, Bozkurt B, Brindis RG, Curtis LH, DeMets D, Hochman JS, Kovacs RJ, Ohman EM, Pressler SJ, Sellke FW, Shen WK, Smith SC, Jr., Tomaselli GF, American College of Cardiology/American Heart Association Task Force on Practice G. 2013 ACC/AHA guideline on the assessment of cardiovascular risk: a report of the American College of Cardiology/American Heart Association Task Force on Practice Guidelines. Circulation 2014;129(25 Suppl 2):S49-73.

34.	Piepoli MF, Hoes AW, Agewall S, Albus C, Brotons C, Catapano AL, Cooney MT, Corra U, Cosyns B, Deaton C, Graham I, Hall MS, Hobbs FDR, Lochen ML, Lollgen H, Marques-Vidal P, Perk J, Prescott E, Redon J, Richter DJ, Sattar N, Smulders Y, Tiberi M, van der Worp HB, van Dis I, Verschuren WMM, Binno S, Group ESCSD. 2016 European Guidelines on cardiovascular disease prevention in clinical practice: The Sixth Joint Task Force of the European Society of Cardiology and Other Societies on Cardiovascular Disease Prevention in Clinical Practice (constituted by representatives of 10 societies and by invited experts)Developed with the special contribution of the European Association for Cardiovascular Prevention & Rehabilitation (EACPR). Eur Heart J 2016;37(29):2315-2381.

35.	Ghiadoni L, Taddei S, Virdis A. Hypertension and endothelial dysfunction: therapeutic approach. Curr Vasc Pharmacol 2012;10(1):42-60.

36.	Luscher TF, Taddei S, Kaski JC, Jukema JW, Kallend D, Munzel T, Kastelein JJ, Deanfield JE, dal VI. Vascular effects and safety of dalcetrapib in patients with or at risk of coronary heart disease: the dal-VESSEL randomized clinical trial. Eur Heart J 2012;33(7):857-65.

37.	Modena MG, Bonetti L, Coppi F, Bursi F, Rossi R. Prognostic role of reversible endothelial dysfunction in hypertensive postmenopausal women. J Am Coll Cardiol 2002;40(3):505-10.

38.	Kitta Y, Obata JE, Nakamura T, Hirano M, Kodama Y, Fujioka D, Saito Y, Kawabata K, Sano K, Kobayashi T, Yano T, Nakamura K, Kugiyama K. Persistent impairment of endothelial vasomotor function has a negative impact on outcome in patients with coronary artery disease. J Am Coll Cardiol 2009;53(4):323-30.

39.	Takishima I, Nakamura T, Hirano M, Kitta Y, Kobayashi T, Fujioka D, Saito Y, Watanabe K, Watanabe Y, Mishina H, Obata JE, Kawabata K, Tamaru S, Kugiyama K. Predictive value of serial assessment of endothelial function in chronic heart failure. Int J Cardiol 2012;158(3):417-22.

40.	Bruno RM, Stea F, Sicari R, Ghiadoni L, Taddei S, Ungar A, Bonuccelli U, Tognoni G, Cintoli S, Del Turco S, Sbrana S, Gargani L, D'Angelo G, Pratali L, Berardi N, Maffei L, Picano E, Train the Brain C. Vascular Function Is Improved After an Environmental Enrichment Program: The Train the Brain-Mind the Vessel Study. Hypertension 2018;71(6):1218-1225.

41.	Ras RT, Fuchs D, Koppenol WP, Garczarek U, Greyling A, Keicher C, Verhoeven C, Bouzamondo H, Wagner F, Trautwein EA. The effect of a low-fat spread with added plant sterols on vascular function markers: results of the Investigating Vascular Function Effects of Plant Sterols (INVEST) study. Am J Clin Nutr 2015;101(4):733-41.

42.	Green DJ, Eijsvogels T, Bouts YM, Maiorana AJ, Naylor LH, Scholten RR, Spaanderman ME, Pugh CJ, Sprung VS, Schreuder T, Jones H, Cable T, Hopman MT, Thijssen DH. Exercise training and artery function in humans: nonresponse and its relationship to cardiovascular risk factors. Journal of applied physiology 2014;117(4):345-52.

43.	Wang TJ. Assessing the role of circulating, genetic, and imaging biomarkers in cardiovascular risk prediction. Circulation 2011;123(5):551-65.

44.	Pignoli P, Tremoli E, Poli A, Oreste P, Paoletti R. Intimal plus medial thickness of the arterial wall: a direct measurement with ultrasound imaging. Circulation 1986;74(6):1399-406.

45.	Woodman RJ, Playford DA, Watts GF, Cheetham C, Reed C, Taylor RR, Puddey IB, Beilin LJ, Burke V, Mori TA, Green D. Improved analysis of brachial artery ultrasound using a novel edge-detection software system. J Appl Physiol 2001;91(2):929-37.

46.	Gemignani V, Faita F, Ghiadoni L, Poggianti E, Demi M. A system for real-time measurement of the brachial artery diameter in B-mode ultrasound images. IEEE Trans Med Imaging 2007;26(3):393-404.

47.	Tomiyama H, Kohro T, Higashi Y, Takase B, Suzuki T, Ishizu T, Ueda S, Yamazaki T, Furumoto T, Kario K, Inoue T, Koba S, Watanabe K, Takemoto Y, Hano T, Sata M, Ishibashi Y, Node K, Maemura K, Ohya Y, Furukawa T, Ito H, Ikeda H, Yamashina A. Reliability of measurement of endothelial function across multiple institutions and establishment of reference values in Japanese. Atherosclerosis 2015;242(2):433-42.

48.	Pyke KE, Tschakovsky ME. The relationship between shear stress and flow-mediated dilatation: implications for the assessment of endothelial function. J Physiol 2005;568(Pt 2):357-69.

49.	Pyke KE, Tschakovsky ME. Peak vs. total reactive hyperemia: which determines the magnitude of flow-mediated dilation? J Appl Physiol 2007;102(4):1510-9.

50.	Padilla J, Johnson BD, Newcomer SC, Wilhite DP, Mickleborough TD, Fly AD, Mather KJ, Wallace JP. Normalization of flow-mediated dilation to shear stress area under the curve eliminates the impact of variable hyperemic stimulus. Cardiovasc Ultrasound 2008;6:44.

51.	Kremkau F. Diagnostic Ultrasound Principles and Instruments: Saunders; 2002.

52.	Thrush A, Hartshorne T. Peripheral Vascular Ultrasound. How, why and when: Elsevier; 2009.

53.	Oates C. Cardiovascular Haemodynamics and Doppler waveforms Explained: Cambridge University Press; 2001.

54.	Boot CR, Groothuis JT, Van Langen H, Hopman MT. Shear stress levels in paralyzed legs of spinal cord-injured individuals with and without nerve degeneration. J Appl Physiol 2002;92(6):2335-40.

55.	Gnasso A, Carallo C, Irace C, Spagnuolo V, De Novara G, Mattioli PL, Pujia A. Association between intima-media thickness and wall shear stress in common carotid arteries in healthy male subjects. Circulation 1996;94(12):3257-62.

56.	Parker BA, Trehearn TL, Meendering JR. Pick your poiseuille: normalizing the shear stimulus in studies of flow-mediated dilation. J Appl Physiol 2009;107(4):1357-9.

57.	Evans DH, Schlindwein FS, Levene MI. The relationship between time averaged intensity weighted mean velocity, and time averaged maximum velocity in neonatal cerebral arteries. Ultrasound Med Biol 1989;15(5):429-35.

58.	Donald AE, Halcox JP, Charakida M, Storry C, Wallace SM, Cole TJ, Friberg P, Deanfield JE. Methodological approaches to optimize reproducibility and power in clinical studies of flow-mediated dilation. J Am Coll Cardiol 2008;51(20):1959-64.

59.	Corretti MC, Anderson TJ, Benjamin EJ, Celermajer D, Charbonneau F, Creager MA, Deanfield J, Drexler H, Gerhard-Herman M, Herrington D, Vallance P, Vita J, Vogel R. Guidelines for the ultrasound assessment of endothelial-dependent flow-mediated vasodilation of the brachial artery: a report of the International Brachial Artery Reactivity Task Force. J Am Coll Cardiol 2002;39(2):257-65.

60.	Greyling A, van Mil AC, Zock PL, Green DJ, Ghiadoni L, Thijssen DH, Dilation TIWGoFM. Adherence to guidelines strongly improves reproducibility of brachial artery flow-mediated dilation. Atherosclerosis 2016;248:196-202.

61.	Thijssen DH, Black MA, Pyke KE, Padilla J, Atkinson G, Harris RA, Parker B, Widlansky ME, Tschakovsky ME, Green DJ. Assessment of flow-mediated dilation in humans: a methodological and physiological guideline. Am J Physiol Heart Circ Physiol 2011;300(1):H2-12.

62.	Dawson EA, Green DJ, Cable NT, Thijssen DH. Effects of acute exercise on flow mediated dilatation (FMD) in healthy humans. Journal of applied physiology (In Press) 2013.

63.	Papamichael CM, Aznaouridis KA, Karatzis EN, Karatzi KN, Stamatelopoulos KS, Vamvakou G, Lekakis JP, Mavrikakis ME. Effect of coffee on endothelial function in healthy subjects: the role of caffeine. Clin Sci (Lond) 2005;109(1):55-60.

64.	Hijmering ML, Stroes ES, Olijhoek J, Hutten BA, Blankestijn PJ, Rabelink TJ. Sympathetic activation markedly reduces endothelium-dependent, flow-mediated vasodilation. J Am Coll Cardiol 2002;39(4):683-8.

65.	Thijssen DH, de Groot P, Kooijman M, Smits P, Hopman MT. Sympathetic nervous system contributes to the age-related impairment of flow-mediated dilation of the superficial femoral artery. Am J Physiol Heart Circ Physiol 2006;291(6):H3122-9.

66.	Ghiadoni L, Donald AE, Cropley M, Mullen MJ, Oakley G, Taylor M, O'Connor G, Betteridge J, Klein N, Steptoe A, Deanfield JE. Mental stress induces transient endothelial dysfunction in humans. Circulation 2000;102(20):2473-8.

67.	Williams MR, Westerman RA, Kingwell BA, Paige J, Blombery PA, Sudhir K, Komesaroff PA. Variations in endothelial function and arterial compliance during the menstrual cycle. J Clin Endocrinol Metab 2001;86(11):5389-95.

68.	ter Avest E, Holewijn S, Stalenhoef AF, de Graaf J. Variation in non-invasive measurements of vascular function in healthy volunteers during daytime. Clin Sci (Lond) 2005;108(5):425-31.

69.	Donald AE, Charakida M, Falaschetti E, Lawlor DA, Halcox JP, Golding J, Hingorani AD, Smith GD, Deanfield JE. Determinants of vascular phenotype in a large childhood population: the Avon Longitudinal Study of Parents and Children (ALSPAC). Eur Heart J 2010;31(12):1502-10.

70.	Widlansky ME, Vita JA, Keyes MJ, Larson MG, Hamburg NM, Levy D, Mitchell GF, Osypiuk EW, Vasan RS, Benjamin EJ. Relation of season and temperature to endothelium-dependent flow-mediated vasodilation in subjects without clinical evidence of cardiovascular disease (from the Framingham Heart Study). Am J Cardiol 2007;100(3):518-23.

71.	Briet M, Collin C, Laurent S, Tan A, Azizi M, Agharazii M, Jeunemaitre X, Alhenc-Gelas F, Boutouyrie P. Endothelial function and chronic exposure to air pollution in normal male subjects. Hypertension 2007;50(5):970-6.

72.	Naylor LH, Weisbrod CJ, O'Driscoll G, Green DJ. Measuring peripheral resistance and conduit arterial structure in humans using Doppler ultrasound. J Appl Physiol 2005;98(6):2311-5.

73.	Doshi SN, Naka KK, Payne N, Jones CJ, Ashton M, Lewis MJ, Goodfellow J. Flow-mediated dilatation following wrist and upper arm occlusion in humans: the contribution of nitric oxide. Clin Sci (Lond) 2001;101(6):629-35.

74.	Green DJ, Dawson EA, Groenewoud HM, Jones H, Thijssen DH. Is flow-mediated dilation nitric oxide mediated?: A meta-analysis. Hypertension 2014;63(2):376-82.

75.	Mullen MJ, Kharbanda RK, Cross J, Donald AE, Taylor M, Vallance P, Deanfield JE, MacAllister RJ. Heterogenous nature of flow-mediated dilatation in human conduit arteries in vivo: relevance to endothelial dysfunction in hypercholesterolemia. Circ Res 2001;88(2):145-51.

76.	Berry KL, Skyrme-Jones RA, Meredith IT. Occlusion cuff position is an important determinant of the time course and magnitude of human brachial artery flow-mediated dilation. Clin Sci (Lond) 2000;99(4):261-7.

77.	Gori T, Dragoni S, Lisi M, Di Stolfo G, Sonnati S, Fineschi M, Parker JD. Conduit artery constriction mediated by low flow. A novel non-invasive method for the assessment of vascular function. J Am Coll Cardiol 2008;51(20):1953-8.

78.	Dawson EA, Alkarmi A, Thijssen DH, Rathore S, Marsman DE, Cable NT, Wright DJ, Green DJ. Low-Flow Mediated Constriction is Endothelium-Dependent: Effects of Exercise Training After Radial Artery Catheterization. Circulation. Cardiovascular interventions 2012;5(5):713-9.

79.	Gori T, Muxel S, Damaske A, Radmacher MC, Fasola F, Schaefer S, Schulz A, Jabs A, Parker JD, Munzel T. Endothelial function assessment: flow-mediated dilation and constriction provide different and complementary information on the presence of coronary artery disease. Eur Heart J 2012;33(3):363-71.

80.	Weissgerber TL, Davies GA, Tschakovsky ME. Low flow-mediated constriction occurs in the radial but not the brachial artery in healthy pregnant and nonpregnant women. J Appl Physiol 2010;108(5):1097-105.

81.	Parker BA, Ridout SJ, Proctor DN. Age and flow-mediated dilation: a comparison of dilatory responsiveness in the brachial and popliteal arteries. Am J Physiol Heart Circ Physiol 2006;291(6):H3043-H3049.

82.	Thijssen DH, van Bemmel MM, Bullens LM, Dawson EA, Hopkins ND, Tinken TM, Black MA, Hopman MT, Cable NT, Green DJ. The impact of baseline diameter on flow-mediated dilation differs in young and older humans. Am J Physiol Heart Circ Physiol 2008;295(4):H1594-8.

83.	Thijssen DH, van Bemmel M, Bullens L, Dawson EA, Hopkins ND, Tinken TM, Black MA, Hopman MT, Cable NT, Green DJ. The impact of baseline diameter on flow mediated dilation (FMD) differs in young and older humans. Am J Physiol Heart Circ Physiol 2008;295(4)(4):H1594-8.

84.	Atkinson G, Batterham AM. Allometric scaling of diameter change in the original flow-mediated dilation protocol. Atherosclerosis 2013;226(2):425-7.

85.	Silber HA, Ouyang P, Bluemke DA, Gupta SN, Foo TK, Lima JA. Why is flow-mediated dilation dependent on arterial size? Assessment of the shear stimulus using phase-contrast magnetic resonance imaging. Am J Physiol Heart Circ Physiol 2005;288(2):H822-8.

86.	Yeboah J, Crouse JR, Hsu FC, Burke GL, Herrington DM. Brachial flow-mediated dilation predicts incident cardiovascular events in older adults: the Cardiovascular Health Study. Circulation 2007;115(18):2390-7.

87.	van den Munckhof I, Scholten R, Cable NT, Hopman MT, Green DJ, Thijssen DH. Impact of age and sex on carotid and peripheral arterial wall thickness in humans. Acta physiologica 2012;206(4):220-8.

88.	Atkinson G, Batterham AM, Thijssen DH, Green DJ. A new approach to improve the specificity of flow-mediated dilation for indicating endothelial function in cardiovascular research. J Hypertens 2013;31(2):287-91.

89.	Leeson P, Thorne S, Donald A, Mullen M, Clarkson P, Deanfield J. Non-invasive measurement of endothelial function: effect on brachial artery dilatation of graded endothelial dependent and independent stimuli. Heart 1997;78(1):22-7.

90.	Black MA, Cable NT, Thijssen DH, Green DJ. Importance of measuring the time course of flow-mediated dilatation in humans. Hypertension 2008;51(2):203-10.

91.	Liuni A, Luca MC, Lisi M, Dragoni S, Di Stolfo G, Mariani JA, Uxa A, Gori T, Parker JD. Observations of time-based measures of flow-mediated dilation of forearm conduit arteries: Implications for the accurate assessment of endothelial function. Am J Physiol Heart Circ Physiol 2010.

92.	Padilla J, Johnson BD, Newcomer SC, Wilhite DP, Mickleborough TD, Fly AD, Mather KJ, Wallace JP. Adjusting flow-mediated dilation for shear stress stimulus allows demonstration of endothelial dysfunction in a population with moderate cardiovascular risk. J Vasc Res 2009;46(6):592-600.

93.	Thijssen DH, Tinken TM, Hopkins N, Dawson EA, Cable NT, Green DJ. The impact of exercise training on the diameter dilator response to forearm ischemia in healthy men. Acta Physiol (Oxf) 2011.

94.	Laughlin MH, Turk JR, Schrage WG, Woodman CR, Price EM. Influence of coronary artery diameter on eNOS protein content. Am J Physiol Heart Circ Physiol 2003;284(4):H1307-12.

95.	Thijssen DH, Dawson EA, Black MA, Hopman MT, Cable NT, Green DJ. Heterogeneity in conduit artery function in humans: impact of arterial size. Am J Physiol Heart Circ Physiol 2008;295(5):H1927-34.

96.	Gemignani V, Bianchini E, Faita F, Giannarelli C, Plantinga Y, Ghiadoni L, Demi M. Ultrasound measurement of the brachial artery flow-mediated dilation without ECG gating. Ultrasound Med Biol 2008;34(3):385-91.

97.	Kizhakekuttu TJ, Gutterman DD, Phillips SA, Jurva JW, Arthur EI, Das E, Widlansky ME. Measuring FMD in the brachial artery: how important is QRS gating? Journal of Applied Physiology 2010;109(4):959-65.

98.	Carter HH, Dawson EA, Birk GK, Spence AL, Naylor LH, Cable NT, Thijssen DH, Green DJ. Effect of SR Manipulation on Conduit Artery Dilation in Humans. Hypertension 2013;61(1):143-50.

99.	Pyke KE, Dwyer EM, Tschakovsky ME. Impact of controlling shear rate on flow-mediated dilation responses in the brachial artery of humans. J Appl Physiol 2004;97(2):499-508.

100.	Anderson EA, Mark AL. Flow-mediated and reflex changes in large peripheral artery tone in humans. Circulation 1989;79(1):93-100.

101.	Furchgott RF, Zawadzki JV. The obligatory role of endothelial cells in the relaxation of arterial smooth muscle by acetylcholine. Nature 1980;288(5789):373-6.

102.	de Groot PC, Poelkens F, Kooijman M, Hopman MT. Preserved flow-mediated dilation in the inactive legs of spinal cord-injured individuals. Am J Physiol Heart Circ Physiol 2004;287(1):H374-80.

103.	Atkinson G, Batterham AM, Black MA, Cable NT, Hopkins ND, Dawson EA, Thijssen DH, Jones H, Tinken TM, Green DJ. Is the ratio of flow-mediated dilation and shear rate a statistically sound approach to normalization in cross-sectional studies on endothelial function? J Appl Physiol 2009;107(6):1893-9.

104.	Betik AC, Luckham VB, Hughson RL. Flow-mediated dilation in human brachial artery after different circulatory occlusion conditions. Am. J. Physiol. 2004;286:442-8.

105.	Thijssen DH, Bullens LM, van Bemmel MM, Dawson EA, Hopkins N, Tinken TM, Black MA, Hopman MT, Cable NT, Green DJ. Does arterial shear explain the magnitude of flow-mediated dilation?: a comparison between young and older humans. Am J Physiol Heart Circ Physiol 2009;296(1):H57-64.

106.	Harris RA, Padilla J. Proper "normalization" of flow-mediated dilation for shear. J Appl Physiol 2007;103(3):1108; author reply 1109.

107.	Atkinson G, Batterham AM, Black MA, Cable NT, Hopkins ND, Dawson EA, Thijssen DH, Jones H, Tinken TM, Green DJ. Is the ratio of flow-mediated dilation and shear rate a statistically sound approach to normalization in cross-sectional studies on endothelial function? Journal of Applied Physiology 2009;107(6):1893-9.

108.	Atkinson G, Batterham AM. The percentage flow-mediated dilation index: a large-sample investigation of its appropriateness, potential for bias and causal nexus in vascular medicine. Vasc Med 2013;18(6):354-65.

109.	Atkinson G. Shear rate normalization is not essential for removing the dependency of flow-mediated dilation on baseline artery diameter: past research revisited. Physiol Meas 2014;35(9):1825-35.

110.	Mitchell GF, Parise H, Vita JA, Larson MG, Warner E, Keaney JF, Jr., Keyes MJ, Levy D, Vasan RS, Benjamin EJ. Local shear stress and brachial artery flow-mediated dilation: the Framingham Heart Study. Hypertension 2004;44(2):134-9.

111.	Philpott AC, Lonn E, Title LM, Verma S, Buithieu J, Charbonneau F, Anderson TJ. Comparison of new measures of vascular function to flow mediated dilatation as a measure of cardiovascular risk factors. Am J Cardiol 2009;103(11):1610-5.

112.	Huang AL, Silver AE, Shvenke E, Schopfer DW, Jahangir E, Titas MA, Shpilman A, Menzoian JO, Watkins MT, Raffetto JD, Gibbons G, Woodson J, Shaw PM, Dhadly M, Eberhardt RT, Keaney JF, Jr., Gokce N, Vita JA. Predictive value of reactive hyperemia for cardiovascular events in patients with peripheral arterial disease undergoing vascular surgery. Arterioscler Thromb Vasc Biol 2007;27(10):2113-9.

113.	Paine NJ, Hinderliter AL, Blumenthal JA, Adams KF, Jr., Sueta CA, Chang PP, O'Connor CM, Sherwood A. Reactive hyperemia is associated with adverse clinical outcomes in heart failure. Am Heart J 2016;178:108-14.

114.	Celermajer DS, Sorensen KE, Bull C, Robinson J, Deanfield JE. Endothelium-dependent dilation in the systemic arteries of asymptomatic subjects relates to coronary risk factors and their interaction. Journal of the American College of Cardiology 1994;24(6):1468-74.

115.	Woodman RJ, Playford DA, Watts GF, Cheetham C, Reed C, Taylor RR, Puddey IB, Beilin LJ, Burke V, Mori TA, Green D. Improved analysis of brachial artery ultrasound using a novel edge-detection software system. Journal of Applied Physiology 2001;91(2):929-37.

116.	Williamson EB, Bronas UG, Dengel DR. Automated edge detection versus manual edge measurement in analysis of brachial artery reactivity: a comparison study. Ultrasound Med Biol 2008;34(9):1499-503.

117.	Harris RA, Nishiyama SK, Wray DW, Richardson RS. Ultrasound assessment of flow-mediated dilation. Hypertension 2010;55(5):1075-85.

118.	Mancini GB, Yeoh E, Abbott D, Chan S. Validation of an automated method for assessing brachial artery endothelial dysfunction. Canadian Journal of Cardiology 2002;18(3):259-62.

119.	Preik M, Lauer T, Heiss C, Tabery S, Strauer BE, Kelm M. Automated ultrasonic measurement of human arteries for the determination of endothelial function. Ultraschall Med 2000;21(5):195-8.

120.	Sonka M, Liang W, Lauer RM. Automated analysis of brachial ultrasound image sequences: early detection of cardiovascular disease via surrogates of endothelial function. IEEE Trans Med Imaging 2002;21(10):1271-9.

121.	Steinbuch J, Hoeks AP, Hermeling E, Truijman MT, Schreuder FH, Mess WH. Standard B-Mode Ultrasound Measures Local Carotid Artery Characteristics as Reliably as Radiofrequency Phase Tracking in Symptomatic Carotid Artery Patients. Ultrasound Med Biol 2016;42(2):586-95.

122.	Faita F, Masi S, Loukogeorgakis S, Gemignani V, Okorie M, Bianchini E, Charakida M, Demi M, Ghiadoni L, Deanfield JE. Comparison of two automatic methods for the assessment of brachial artery flow-mediated dilation. J Hypertens 2011;29(1):85-90.

123.	Greyling A, van Mil AC, Zock PL, Green DJ, Ghiadoni L, Thijssen DH, Dilation TIWGoFM. Assessing the perceived quality of brachial artery Flow Mediated Dilation studies for inclusion in meta-analyses and systematic reviews: Description of data employed in the development of a scoring ;tool based on currently accepted guidelines. Data Brief 2016;8:73-7.

124.	Craiem D, Chironi G, Gariepy J, Miranda-Lacet J, Levenson J, Simon A. New monitoring software for larger clinical application of brachial artery flow-mediated vasodilatation measurements. J Hypertens 2007;25(1):133-40.

125.	Herrington DM, Fan L, Drum M, Riley WA, Pusser BE, Crouse JR, Burke GL, McBurnie MA, Morgan TM, Espeland MA. Brachial flow-mediated vasodilator responses in population-based research: methods, reproducibility and effects of age, gender and baseline diameter. J Cardiovasc Risk 2001;8(5):319-28.

126.	Simova I, Nossikoff A, Denchev S. Interobserver and intraobserver variability of flow-mediated vasodilatation of the brachial artery. Echocardiography 2008;25(1):77-83.

127.	Witte DR, van der Graaf Y, Grobbee DE, Bots ML, Group SS. Measurement of flow-mediated dilatation of the brachial artery is affected by local elastic vessel wall properties in high-risk patients. Atherosclerosis 2005;182(2):323-30.

128.	Vanoli D, Wiklund U, Lindqvist P, Henein M, Naslund U. Successful novice's training in obtaining accurate assessment of carotid IMT using an automated ultrasound system. Eur Heart J Cardiovasc Imaging 2014;15(6):637-42.

129.	Bruno RM, Ghiadoni L. Vascular smooth muscle function: defining the diabetic vascular phenotype. Diabetologia 2013;56(10):2107-9.

130.	Joannides R, Haefeli WE, Linder L, Richard V, Bakkali EH, Thuillez C, Luscher TF. Nitric oxide is responsible for flow-dependent dilatation of human peripheral conduit arteries in vivo. Circulation 1995;91(5):1314-9.

131.	Joannides R, Richard V, Haefeli WE, Benoist A, Linder L, Luscher TF, Thuillez C. Role of nitric oxide in the regulation of the mechanical properties of peripheral conduit arteries in humans. Hypertension 1997;30(6):1465-70.

132.	Maruhashi T, Soga J, Fujimura N, Idei N, Mikami S, Iwamoto Y, Kajikawa M, Matsumoto T, Hidaka T, Kihara Y, Chayama K, Noma K, Nakashima A, Goto C, Higashi Y. Nitroglycerine-induced vasodilation for assessment of vascular function: a comparison with flow-mediated vasodilation. Arterioscler Thromb Vasc Biol 2013;33(6):1401-8.

133.	Montero D, Walther G, Perez-Martin A, Vicente-Salar N, Roche E, Vinet A. Vascular smooth muscle function in type 2 diabetes mellitus: a systematic review and meta-analysis. Diabetologia 2013;56(10):2122-33.

134.	Fontes-Guerra PC, Cardoso CR, Muxfeldt ES, Salles GF. Nitroglycerin-mediated, but not flow-mediated vasodilation, is associated with blunted nocturnal blood pressure fall in patients with resistant hypertension. J Hypertens 2015;33(8):1666-75.

135.	Kullo IJ, Malik AR, Bielak LF, Sheedy PF, 2nd, Turner ST, Peyser PA. Brachial artery diameter and vasodilator response to nitroglycerine, but not flow-mediated dilatation, are associated with the presence and quantity of coronary artery calcium in asymptomatic adults. Clin Sci (Lond) 2007;112(3):175-82.

136.	Malik AR, Sultan S, Turner ST, Kullo IJ. Urinary albumin excretion is associated with impaired flow- and nitroglycerin-mediated brachial artery dilatation in hypertensive adults. J Hum Hypertens 2007;21(3):231-8.

137.	Morimoto S, Yurugi T, Aota Y, Sakuma T, Jo F, Nishikawa M, Iwasaka T, Maki K. Prognostic significance of ankle-brachial index, brachial-ankle pulse wave velocity, flow-mediated dilation, and nitroglycerin-mediated dilation in end-stage renal disease. Am J Nephrol 2009;30(1):55-63.

138.	Akamatsu D, Sato A, Goto H, Watanabe T, Hashimoto M, Shimizu T, Sugawara H, Sato H, Nakano Y, Miura T, Zukeran T, Serizawa F, Hamada Y, Tsuchida K, Tsuji I, Satomi S. Nitroglycerin-mediated vasodilatation of the brachial artery may predict long-term cardiovascular events irrespective of the presence of atherosclerotic disease. J Atheroscler Thromb 2010;17(12):1266-74.

139.	Ducharme A, Dupuis J, McNicoll S, Harel F, Tardif JC. Comparison of nitroglycerin lingual spray and sublingual tablet on time of onset and duration of brachial artery vasodilation in normal subjects. Am J Cardiol 1999;84(8):952-4, A8.

140.	Ghiadoni L, Huang Y, Magagna A, Buralli S, Taddei S, Salvetti A. Effect of acute blood pressure reduction on endothelial function in the brachial artery of patients with essential hypertension. J Hypertens 2001;19(3 Pt 2):547-51.

141.	Deanfield J, Donald A, Ferri C, Giannattasio C, Halcox J, Halligan S, Lerman A, Mancia G, Oliver JJ, Pessina AC, Rizzoni D, Rossi GP, Salvetti A, Schiffrin EL, Taddei S, Webb DJ. Endothelial function and dysfunction. Part I: Methodological issues for assessment in the different vascular beds: a statement by the Working Group on Endothelin and Endothelial Factors of the European Society of Hypertension. J Hypertens 2005;23(1):7-17.

142.	Oliver JJ, Bowler A, Beudeker Q, Cate T, Webb DJ. Dose-response relationship of sublingual nitroglycerin with brachial artery dilatation and change in central and peripheral augmentation index. Clinical pharmacology and therapeutics 2005;77(4):337-8.

143.	Ghiadoni L, Taddei S, Bruno RM. Hemodynamic and autonomic effects of low-dose glyceryl trinitrate used to test endothelium-independent vasodilation of the brachial artery. Artery Research Volume 2017;20:85.

144.	Bressler B, Chan S, Mancini GB. Temporal response of brachial artery dilation after occlusion and nitroglycerin. Am J Cardiol 2000;85(3):396-400, A10.

145.	Thelen AM, Kelly AS, Williamson EB, Dengel DR. Examining the time course of endothelium-independent dilation by nitroglycerin. Ultrasound Med Biol 2008;34(8):1217-20.

FIGURE LEGENDS

FIGURE 1. A: Schematic presentation of diameter and shear stress (or rate) responses before and after a 5-minute ischaemic stimulus applied distally to the forearm.

B: representative diameter (upper panel) and velocity (lower panel) data from a healthy person before and after the 5-minute ischaemic stimulus.

C: representative diameter (upper panel) and velocity (lower panel) data from a person with endothelial dysfunction before and after the 5-minute ischaemic stimulus.

FIGURE 2. Experimental set-up for the performance of the brachial artery FMD.

A: Position of the cuff and ultrasound probe placed on the upper arm.

B: Screen shot of a representative B-mode image of the diameter and Doppler blood flow velocity.

C: Wall tracking software to perform (semi)automated analysis.

D: Output generated by the analysis software to enable calculation of the FMD.